## Slide 1
Power Supply Box
120V AC
5V
12V
24V
28V
Salinity Sensor Driver Box
5V
UART
Speedgoat
Digital Data
Power
Analog Signal
Pressure and Flowrate Sensors Passthrough
12V
28V
0-10V (Press.)
Pulse (flowrate)
UART1
A I/O
D I/O
UART2
Control Valve Driver Box
12V
UART
Primary Interconnections
24V
A I/O
5V, 12V, 24V, 28V (reference)

## Slide 2
Power Supply Box
120V AC
12V
24V
28V
5.5mm x 2.1mm barrel jack sockets
Linear Voltage Regulated Board
5V
12V
24V
DB9_0
1
2
3
4
5-9
DC Ground
DB9_0 pinout:
5V DC (1/3 gain)
12V DC (1/3 gain)
24V DC (1/3 gain)
28V DC (1/3 gain)
5-9. DC ground
Linear Power Supply
28V
28V
28V
Banana plug sockets
1/3
1/3
1/3
1/3
5V

## Slide 3
Power Supply Box
120V AC
5V
12V
24V
28V
5.5mm x 2.1mm barrel jack sockets
Meanwell RT-50D PSU
120V
5V
12V
24V
Mad Hornets 12V  28V Converter
12V
28V
DB9_0
1
2
3
4
5-9
1/3
1/3
1/3
DC Ground
DB9_0 pinout:
5V DC (1/3 gain)
12V DC (1/3 gain)
24V DC (1/3 gain)
28V DC (1/3 gain)
5-9. DC ground
1/3

## Slide 5
Salinity Sensor Driver Box
EZO Cond. Driver Board
5V
I2C
EZO Temp. Driver Board
5V
I2C
EZO Cond. Driver Board
5V
I2C
EZO Temp. Driver Board
5V
I2C
EZO Cond. Driver Board
5V
I2C
EZO Temp. Driver Board
5V
I2C
SMA screw connectors
5V
Breadboard
Arduino
UART
UART (USB)
Speedgoat
PSU Box
I2C
Sensor 1
Sensor 2
Sensor 3
Digital Data
Power

## Slide 7
Pressure and Flowrate Sensors Passthrough
28V
PSU Box
12V
PSU Box
5.5mm x 2.1mm barrel jack sockets
DB9_1
Protoboard
1
2
3
4
5
6
4 pin waterproof AMP connectors
Press 1
Press 2
Flow 1
Flow 2
Flow 3
Flow 4
DB9_1 pinout:
0-10V, Press. 1
0-10V, Press. 2
0-10V, Press. 3
5V Pulse, Flow 1
5V Pulse, Flow 2
5V Pulse, Flow 3
5V Pulse, Flow 4
0-10V, Flow 5
Ground
Flow 5
Press 3
7
8
9

## Slide 8
Control Valve Driver Box
24V
PSU Box
12V
PSU Box
5.5mm x 2.1mm barrel jack sockets
NCD 4-20 mA I2C transmitter
Breadboard
Arduino Uno
w/ I2C Hat
UART
I2C
UART (USB)
Speedgoat

## Slide 9
Speedgoat Breakout Box
DB9_1
2
3
4
5
6
7
8
9
1
DB9_0
2
3
4
5
1
DB9_1 pinout:
0-10V, Press. 1
0-10V, Press. 2
0-10V, Press. 3
5V Pulse, Flow 1
5V Pulse, Flow 2
5V Pulse, Flow 3
5V Pulse, Flow 4
0-10V, Flow 5
Ground
DB9_0 pinout:
5V DC (1/3 gain)
12V DC (1/3 gain)
24V DC (1/3 gain)
28V DC (1/3 gain)
5-9. DC ground
*Connect IO397 2a, 4a, 6a, and 8a to analog ground (13a, 14a, or 17a).
^AI1-4 and A01-4 are RCA sockets. Connect ground to analog ground bus bar. Unused analog pins must be connected to analog ground.
1a*
3a*
5a*
7a*
13a
1a
2a
3a
4a
1b
2b
3b
4b
17a
17a
5a
6a
7a
8a
Analog Ground Bus Bar
AI1^
AI2^
AI3^
AI4^
AO1^
AO2^
AO3^
AO4^
9a
10a
11a
12a
IO397
IO191

## Slide 10
DB9_1 pinout:
0-10V, Press. 1  IO191, 1a
0-10V, Press. 2  IO191, 2a
0-10V, Press. 3  IO191, 3a
5V Pulse, Flow 1  IO191, 1b
5V Pulse, Flow 2  IO191, 2b
5V Pulse, Flow 3  IO191, 3b
5V Pulse, Flow 4  IO191, 4b
0-10V, Flow 5  IO191, 4a
Ground  IO191, 13a
DB9_0 pinout:
5V DC (1/3 gain)  IO397, 1a
12V DC (1/3 gain)  IO397, 3a
24V DC (1/3 gain)  IO397, 5a
28V DC (1/3 gain)  IO397, 7a
5-9. DC ground  IO397, 13a
IO191 mapping:
1a: Pressure sensor 1, 0-10V
2a: Pressure sensor 2, 0-10V
3a: Pressure sensor 3, 0-10V
4a: Flowrate sensor 5, 0-10V
5a: External analog input 1 (load cell)
6a: External analog input 2 (LVDT)
7a: External analog input 3 (ground unless used)
8a: External analog input 4 (ground unless used)
9a: External analog output 1 (DLE control valve)
10a: External analog output 2
11a: External analog output 3
12a: External analog output 4
13a: DB9_1 ground
14a: Connect to IO191 pin 17b
17a: Connect to IO397 pin 17a
1b: Flowrate sensor 1, 5V pulse
2b: Flowrate sensor 2, 5V pulse
3b: Flowrate sensor 3, 5V pulse
4b: Flowrate sensor 4, 5V pulse
17b: Connect to IO191 pin 14a
IO397 mapping:
1a: 5V DC reference (1/3 gain)
2a: 12V DC reference (1/3 gain)
3a: 24V DC reference (1/3 gain)
4a: 28V DC reference (1/3 gain)
13a: DB9_0 ground
17a: Connect to IO191 pin 17a

## Slide 11
Shopping list:
7-pin JST sockets
DB9 sockets
Four-conductor wire
Protoboards
3x Teensy 4.1
6-32 x 0.25” heat-set inserts
6-32 SHC screws
Crimp butt connectors
Soldering iron tips